## Slide 1: Polychlorotrifluoroethylene (PCTFE) Polymer Enclosure for an Implantable Device
Neuralink, US 2023/0154809 A1, Published 5/18/2023
6/13/2024

## Slide 2: Background
Implantable devices in-vivo have rigid hermetic housing to protect active electronics from environmental and mechanical stress
Increasingly common for large number (e.g., thousands) of tightly-spaced wires extending from implantable device
Traditional housing materials can become problematic with above structure (e.g., cracked glass housing)

## Slide 3: Outline
Goal: Describe a polychlorotrifluoroethylene (PCTFE) polymer enclosure with tightly spaced wires extending out of it
Describe PCTFE + Why It’s Used
Embodiments:
  Implantable Device with PCTFE Enclosure + Exploded View
  5 Views of a PCTFE Enclosure with Insulated Wires
  Sealing System for Manufacturing a PCTFE Enclosure
  2 Views of an Insert of a Sealing System for Manufacturing a PCTFE Enclosure
  Sealing System in Operation for Manufacturing a PCTFE Enclosure
  Controller for Use with the Sealing System
  Flowchart for Manufacturing Implantable Device
  Computer System

## Slide 4: What is PCTFE?
Thermoplastic chloro-fluoropolymer (), where n – number of monomer units in polymer molecule
Manufactured and sold under trade names like KelF, FluoroPro, Neoflon, Aclar

## Slide 5: Why PCTFE?
Chemical properties:
  Good resistance to chemical attack and oxidizing agents – high fluorene content, lack of hydrogen atoms
Physical properties:
  Good rigidity to prevent mechanical stress on implantable device
  Low dielectric constant for Bluetooth Low Energy (BLE) communication and wireless charging
Barrier properties:
  Close to zero water absorption (<0.01%)
  Good for substantially hermetic enclosure (completely prevent passage of fluids/gases for tens of years)
  Low water vapor permeability
  Highest gas barrier property with WVTR (0.005g/m2/24 hr) at 25C
Mechanical properties:
  High tensile strength
  Moderate hardness
Thermal properties:
  Melting point: 212C
  Maximum temperature for continuous use: 120C
  Glass transition temperature: 45C
  Good melt flow properties with excellent stress crack resistance

## Slide 6: Implantable Device with PCTFE Enclosure
PCTFE Enclosure
First PCTFE portion
Second PCTFE portion
Insulated Wires
(at least 500, parallel)
Welded Seam
(seals 2 PCTFE portions)
Attachment Means
(holds implantable device in place near skull of subject)
Screw
hole:
1.8-1.9
mm
Screws: 1.8mm diameter

## Slide 7: Exploded view of Implantable Device
Attachment Means
PCTFE Enclosure
First portion
Second portion
Insulated Wires

## Slide 8: Close-up Cross-sectional View of a PCTFE Enclosure with Insulated Wires
Welded Seam
(adheres around insulation of wires if PCTFE portions hermetically sealed)
Insulated Wires
- Composed of polymide (PI)
- Thin and long -
Width 5-50 micrometers,
Nominal thickness 4-6 micrometers,
Length about 20mm
- Produced by lithography
Second PCTFE portion
First PCTFE portion

## Slide 9: Close-up Cross-sectional View of a PCTFE Enclosure with Insulated Wires
Welded Seam
First PCTFE portion
Second PCTFE portion
Metal trace
Individual, metallic threads within insulation
Include metals/metal alloys (e.g., thin gold film)
May include 16 traces
Terminates in electrode
Insulated Wire
Distance between metal traces
(2 micrometers)
Cross-sectional width of metal trace
(2 micrometers)

## Slide 10: Closer View of a PCTFE Enclosure with Insulated Wires
Insulated Wire
Metal Traces
Second PCTFE portion
First PCTFE portion

## Slide 11: View of a PCTFE Enclosure with Insulated Wires
Insulated Wire
Distance
(less than 150 micrometers)
PCTFE Enclosure

## Slide 12: (Even Closer) View of a PCTFE Enclosure with Insulated Wires
Insulated Wire
Metal Traces
PCTFE Enclosure

## Slide 13: Figs 1-6 Manufactured via Sealing System
Figs 1-6 manufactured via sealing system
Next figures will describe sealing system + process

## Slide 14: Sealing System for Manufacturing a PCTFE Enclosure
Opening
Base
Hold PCTFE enclosure
Heater
Apply heat to PCTFE enclosure

## Slide 15: First View of an Insert of a Sealing System for Manufacturing a PCTFE Enclosure
PCTFE Enclosure
Wires
Screws
Attachment means
Circuitry
Clamp
Insert
Holds wires and PCTFE enclosure in place during sealing

## Slide 16: Second View of an Insert of a Sealing System for Manufacturing a PCTFE Enclosure
PCTFE Enclosure
Clamp
Wires
Circuitry

## Slide 17: Sealing System in Operation for Manufacturing a PCTFE Enclosure
Base
Insert
PCTFE Enclosure
Insulated Wires
Heater
Pistons
- Apply pressure to PCTFE enclosure via hydraulic/pneumatic cylinders
- Includes 2 pistons (one above base, one below base)
Air Nozzles
Apply air to enclosure during/after sealing process (e.g., cool enclosures after sealing, cool enclosure to keep temperature reasonable during sealing)
PCTFE melts and reflow around insulated wires for sealing under appropriate temperatures

## Slide 18: Controller for Use with the Sealing System
Pressure gauge for top piston
Pressure gauge for bottom piston
Pressure gauge for air
Temperature Controls
Power button
Display
(shows current temperature)
Dial
Controller:
Automatically controls sealing process

## Slide 19: Flowchart for Manufacturing Implantable Device
Sealing process: May be partially or completely automated!

## Slide 20: Step 1:
PCTFE Enclosure
Wires
Screws
Attachment means
Circuitry
Clamp

## Slide 21: Step 2:
Base
Insert
PCTFE Enclosure
Insulated Wires
Heater
Pistons
Air Nozzles
Pressure applied to first piston: 40 – 80 psi
Pressure applied to second piston: > 100 psi

## Slide 22: Step 3:
Opening
Base
Heater
Heater heats PCTFE to at least 200C to melt PCTFE for the sealing process
Pressure applied to the heating for the thermocompression molding process

## Slide 23: Step 4:
1. Thermally melt 1st and 2nd PCTFE portions & adhere around insulation of each wire to seal enclosure
PCTFE melts and reflows around wires
Polymer + insulation – forms a seal around wire
Heat PCTFE to 200-300C
Must set upper limit temperature (e.g., 255-260C) to avoid melting other components
2. Cool Implantable Device
Sealing system transitions to cooling mode
Notify user (e.g., alarm) when cooling complete
User detatches implantable device from sealing system (e.g, via clamps, unscrewing, etc.)

## Slide 24: Computer System