## Slide 1: Out-of-Band Pairing for Wireless Neural Implant
Neuralink
US 2024/0232340 A1
Published 7/11/2024

## Slide 2: Background
Brain-machine interfaces require fine-scale communication at the level of individual neurons - large #s across multiple brain areas
A large amount of information needed for processing & communication to outside world
Wireless solutions needed for more convenient communication & processing
Challenge of wireless solutions – data vulnerable to man-in-the-middle (MTTM) attacks
Need to securely transfer data to and from brain wirelessly

## Slide 3: Outline
Goal: Describe a secure wireless pairing between neural implant & remote device that is via magnetic induction & neural signal analysis
Embodiments:
  Perspective View of System with Implanted Device + External Device
  Sectional Side View of Implanted Device
  Normal Device Pairing
  Attempt of Man-In-The-Middle Attack
  Detection of Man-In-The-Middle Attack via Out-Of-Band Pairing Confirmation
  Enabling & Aborting a Key Exchange via Out-Of-Band Pairing Code
  Flowchart: Secure Pairing via Magnetic Induction
  Flowchart: Secure Pairing via Neural Signal Analysis
  Flowchart: Code Pairing Confirmation
  Example Computer System

## Slide 4: Out-Of-Band Authentication
Secondary verification through separate communication channel
Introduce additional layer of security via multiple, unrelated channels
Example: SMS one-time passwords, push notifications, biometric authentication (facial, voice, neural signals, etc.)

## Slide 5: Perspective View of System with Implanted Device + External Device
Implanted device
Brain
Charging device
Send/receive signals via magnetic induction optically, or by amplitude/frequency modulation
Subject
External device (e.g., phone, laptop, etc.)
Relay device
5-10mm
Note: Charging & External device can be same device

## Slide 6: Sectional Side View of Implanted Device
Circuitry
IC chips for on-board amplification & digitization of data
Includes amplifiers, ADCs, and multiplexing electronics for digital packetization of brain signals
Can perform spike detection via thresholding or voltage vs time-analysis
Flexible electrodes
Biocompatible case
Charging coil
(receive power from varying magnetic field via charging device)
Small battery
Communications component
wireless transceiver
Initiates uplink communication via pairing techniques

## Slide 7: Key Exchange for Normal Device Pairing
External Device
Implanted Device
Key exchange – devices communicate over public channel after secure pairing, involves constructing shared secret
Shared secret – used as encryption key on public channel after key exchange
Public key (external)
Public key (implanted)

## Slide 8: Diffie-Hellman Key Exchange (DHKE)
External device - generate external device public and private keys (PKa and PRIVa)
Implanted device - generate implanted device public and private keys (PKb and PRIVb)
External Device
Implanted Device
Public key (external)
Public key (implanted)
- private key PRIVa of external device
- public key PKa of external device
- private key PRIVb of implanted device
- public key PKb of implanted device
– large prime number

## Slide 9: Attempt of Man-In-The-Middle Attack
Resistant to passive eavesdropping – computing public keys is infeasible without knowing private keys

## Slide 10: Detection of Man-In-The-Middle Attack via Out-Of-Band Pairing Confirmation
Out-of-band check:
Implanted & external devices generate random pairing codes derived from key exchange to third party via separate channel (see Bluetooth specs)
Compare both values (NOT match – MITM attack!)

## Slide 11: Out-Of-Band Data Transport: Enabling & Aborting a Key Exchange
Key exchange
Out-of-Band Data Transport – passkey/pairing code sent one-way from external device to implanted device (different from out-of-band check)
Use commit-reveal MITM protection (see Bluetooth specs)
Exchange hash outputs generated from candidate public keys, random nonce, and single bit of passkey at a time
If single output does not match output expected from any device  abort key exchange

## Slide 12: Secure Pairing via Magnetic Induction from Helper Device

## Slide 13: Secure Pairing via Neural Signal Analysis

## Slide 14: Secure Pairing via Pairing Code Transmission

## Slide 15: Example Computer System