## Slide 1: Dicing and Partial Dicing Release of Lithographed Device
Neuralink, US 2023/0382098, 11/30/2023 6/19/2024

## Slide 2: Background
System and method for implanting devices to biological tissue described in US 2018/0296243 (“Methods, Compositions, and System for Device Implantation”)[1]
System and method for implanting devices to biological tissue described in US 2020/0086111 (“Device Implantation using a Cartridge”)
Before implanting and piercing biological tissue, need to remove electrodes from substrate (expose electrodes for implantation)
[1] This patent is NOT from Neuralink

## Slide 3: Outline
Goal: Describe method for fabricating cantilever structure using partial dicing (partitioning of wafer)
Embodiments:
  Method for fabricating cantilever structure
  Cross-sectional views of structures fabricated with and without sacrificial layer
  Cross-sectional view of substrate
  Top-down views of substrate
  Wafer map

## Slide 4: Method for fabricating cantilever structure

## Slide 5: Step 1:
Substrate (silicon wafer)
Substrate (silicon wafer)
50 – 55 mm
8 – 15 mm
without sacrificial layer
with sacrificial layer

## Slide 6: Step 2:
Sacrificial layer (e.g., alumninum, chromium, silicon oxide, titanium, tungsten)
with sacrificial layer
without sacrificial layer

## Slide 7: Step 3:
Compliant layer (device stack) (e.g., polymide, parylene-C, SU-8 photoresist, benzocyclobutene (BCB))
with sacrificial layer
without sacrificial layer

## Slide 8: Step 4:
with sacrificial layer
without sacrificial layer
Partial cut on BOTTOM – done to minimize detrimental effects from dicing
(e.g., accidental damage to overhang structure, contamination)
Can make partial cuts on sacrificial layer without post-release breaking, but very difficult for thin sacrificial layers; can remove residual substrate from material from partial cut

## Slide 9: Step 5:
with sacrificial layer
without sacrificial layer
Etchant – dilute HF, buffered HF, nitric acid
Die-level processing (e.g., solvent cleans) – release device stack using wet or dry etchant
(release overhang structure AFTER partial cuts – minimize inadvertent damage to device stack & contamination)

## Slide 10: Step 6:
with sacrificial layer
without sacrificial layer
Break-away operation:
- Can be performed in air OR in solution;
- Can be performed manually (tweezers/forceps) or automated (impulse bar & fixturing)

## Slide 11: Cross-sectional views of structures fabricated with sacrificial layer
Steps 4-6 above

## Slide 12: Cross-sectional views of structures fabricated without sacrificial layer
Steps 4-6 above

## Slide 13: Cross-section view of substrate
Silicon wafer substrate
(aka partial cut)
Made from partial saw-and-break method
Die portions MUST be separated to ensure device integrity
Partial saw blade: 50 - 254 microns (10 mils), cut 85-99% of substrate
Full cut: 100% through substrate, 10-mil or smaller saw blade

## Slide 14: Top-down view of substrate
Free space
Exposed electrodes
(exposed to allow robot to implant)

## Slide 15: Top-down view of substrate
Head
(thread management / facilitating implantation;
hold array of threads / electrodes)
Body
(Peel force management)
Pill
(Electrochemical testing)

## Slide 16: Top-down view of substrate
Threads – NOT bent during breaking
Maintain aluminum release layer:
Dice singulate arrays
Etch aluminum
’Snap’/break the neck
 Device (PI) “floating” (NOT on silicon) due to aluminum underneath etched away

## Slide 17: Sample Wafer Map
Enlarged die
(shows more detail)
mm
Wafer map
Wafer map
each