## Slide 1: In Vitro Neural Implant Tester with Hardware-In-The-Loop Simulation
Neuralink, US 2023/0147266 A1, Published 5/11/2023
6/11/2024

## Slide 2: Background
Functionality of the implantable device depends on active electronic devices
Body environments (e.g., brain tissue) have high humidity and bodily fluids
Implantable devices in-vivo (inside organism) environment with high humidity and oxidative inflammatory response puts stress on sensitive components like active electronic devices
Need failure analysis for implantable devices to determine an implantation period
Need to test large number of devices
Need to accurately replicate in-vivo conditions in a testing environment
Implanted devices take years to fail – need more time-efficient testing environment

## Slide 3: Outline
Goal: Describe accelerated testing apparatuses for implants to stimulate the circulation of bodily fluids via performing accelerated aging tests of many wireless neural implants simultaneously
Purpose: To determine the implantation period of each implant via accelerated testing in simulated in-vivo environments for many implants at once
Embodiments:
  Accelerated Testing Apparatus
  Vessel + Fixtures of Accelerated Testing Apparatus
  Closed vessel of Accelerated Testing Appratus
  Fixture with Yoked Parts
  Side View of Fixture
  Front view of Accelerated Testing Apparatus
  Flowchart

## Slide 4: Accelerated Testing Apparatus
Cabinet
(18U or greater size server rack)
Multiple Bays
Watertight basin
RF transparent lid
Vessel
Antenna
Computer server
Latching mechanism
(e.g., clamp, latch)
Force together lid & watertight basin to hermetically seal vessel
Placed on shelf Insertable/removable
3D printed
Modular, “substantially similar”
Holds 200 mL of liquid
- Computer server coupled with antenna for communication with devices-under-test
- Data (e.g., internal physical properties) of devices-under-test continuously received, indexed by computer server
- Can include multiple computer servers

## Slide 5: Front view of Accelerated Testing Apparatus
Pressure Sensor
Reservoir
Peroxide sensor
Peroxide dispenser
Pump
Heater
Metal portions
Controller
Multiple Bays
Cabinet
Vessel
RF transparent lid (open)
Watertight basin
Latching mechanism
Computer server
Antenna
Reservoir:
Includes liquid with properties similar of cerebral spinal fluid (CSF)
Synthetic CSF – chemically mimic in-vivo environment via solvent pH, ionic strength, conductivity, reactive oxygen species
Holds 8L of liquid
Disposed on external surface of reservoir
Measure pressure between reservoir and vessel
Peroxide sensor:
Measures pH, ionic concentration, temperature, or pressure
Peroxide – allow temperature acceleration without unrealistic failure
Peroxide concentration: higher than in-vivo levels, without unrealistic failure (e.g., 2.5 millimolars)
Vessel + Reservoir:
Share common liquid
Pressurized to at least 202,650 pascals (2 atm)
pH of liquid: 7.4
Ionic concentration 6000 microsiemens/cm
Vessel:
Holds 200mL of liquid
Peroxide dispenser:
Connected with computer server, which commands dispensing peroxide from peroxide dispenser (e.g., stop dispensing if max limit reached, start dispensing if min limit reached)
Pump:
Pump liquid from reservoir to vessel
Circulates constantly replenishing liquid, similar to brain fluid (rather than using a heating stationary liquid replaced every couple of days)
Heater:
Heat liquid to above body temperature (e.g., 37 to 57C), just enough to fail realistically
Immersion heater (metal portions along immersion length of immersion heater coated in polymer)
Controller:
Pulse pressure from pump to induce & exaggerate cyclic mechanical fatigue
Maintain pressure within a yield stress to decrease time to failure, without non-reversible effects

## Slide 6: Vessel + Fixtures of Accelerated Testing Apparatus
Vessel
RF transparent lid
Watertight basin
(e.g., )
Latching mechanism
Dock
Power transmission coil
- Powers respective device-under-test within fixture, provide wireless charging
Fixtures
(e.g., 16 fixtures per vessel, each fixture holds device-under-test)
Fixtures

## Slide 7: Closed vessel of Accelerated Testing Apparatus
RF transparent lid
Watertight basin
Latching mechanism
Fixture

## Slide 8: Fixture with Yoked Parts
First part
Second part
Simulated tissue
Device well

## Slide 9: Side view of Fixture
First part
Device well (body)
Simulated tissue well
(electrodes)
Second part
Electrodes – sewn into brain-like material (conductive agarose) to mimic mechanical stresses of electrodes

## Slide 10: Reservoir + Other Components of Accelerated Testing Apparatus
Reservoir
Peroxide sensor
Pressure Sensor
(may measure pH, ionic concentration, temperature of liquid within reservoir)
Heater
Pump
Peroxide dispenser

## Slide 11: Flowchart